Operation Manual of the Special Handset

- Keys on the handset:

Menu, Enter, ESC, UP, DOWN, 1-9 Number buttons.

- How to change the quantity of the column in the system:

- Plug the special handset to the control box. There is JIECANG LOGO on the handset.

- Press “Menu” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

SET  Parameters

Normal Run

(The “Normal Run” is fleshing). Press “UP” button, the “SET Parameters” will flesh. Press “Enter” and it will list:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn001：2 pcs

Motor Quantity 1-4

(The “2” is fleshing. It means the default column quantity for the system is 2 columns). Press “3”, it will flesh “3”. Press “Enter” and the “3” will not flesh. It means you have changed to 3 columns for the system already (the parameter of the program inside the control box has been changed).

3. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn002:CM

1—INCH   0---CM

“CM” is fleshing. If you Press “1”, the “INCH” will flesh. If you press “Inter”, it will not flesh and the parameter is changed now.

4. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn003：30CM

Deacceleration  position

“30” is fleshing. You change “30” as the same as above process.

5. Press “DOWN” buttons, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn004：5A

Reset  current

“5” is fleshing. You can change “5” accordingly.

6. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn005: 650mm

Setting High Position

“650” is fleshing. You can change “650” accordingly.

7. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn006: 5 mm

Setting Low Position

“5” is fleshing. You can change “5” accordingly.

8. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn007: 176

Electric Gear Ratio Steps Per CM

“176” is fleshing. You can change “176” accordingly.

9. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn008: 560 mm

Start Height

“560” is fleshing. You can change “560” accordingly.

10. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn009: 10 A

Motor Overcurrent protection

“10” is fleshing. You can change “10” accordingly.

11. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn010: 8

Overload protection sensitivity

“8” is fleshing. You can change “8” accordingly.

12. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn011: 30 mm

Overload Protection Untread Distance

“30” is fleshing. You can change “30” accordingly.

13. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn012: 75 ℃

Hot protection temperature

30--99

“75” is fleshing. You can change “75” accordingly.

Note: If the number you print is more than 75, it will change to 75 automatically. If you print the number is less than 30, it will change to 30 automatically.

14. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn013: 45 ℃

Release Hot protection temperature

“45” is fleshing. You can change “45” accordingly.

Note: If the number you print is more than 75, it will change to 75 automatically. If you print the number is less than 30, it will change to 30 automatically.

15. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn014：35 mm/s

Constant Speed

“35” is fleshing. You can change “35” accordingly.

16. Press “DOWN” button, it will list following:

JIECANG MEDICAL EQUIPMENT CO.LTD

Pn015：1 S(0-9)

Acceleration Time

“1” is fleshing. You can change “1” accordingly.

Explanation of the Parameters (Table 1)

| Parameters | Explanation | Default |
| --- | --- | --- |
| Pn001 | Motor Quantity | 2 |
| Pn002 | Inch/CM | CM |
| Pn003 | Deacceleration position | 30 |
| Pn004 | Reset Current | 5A |
| Pn005 | Setting High Position | 650mm |
| Pn006 | Setting Low Position | 5m |
| Pn007 | Electric Gear Ratio Steps Per Cm | 176 |
| Pn008 | Start Height | 560mm |
| Pn009 | Motor overcurrent protection | 10A |
| Pn010 | Overload Protection Sensitivity | 8 |
| Pn011 | Overload Protection untread distance | 30mm |
| Pn012 | Hot Protection temperature | 75 |
| Pn013 | Release Hot protection temperature | 45 |
| Pn014 | Constant Speed | 34mm/s |
| Pn015 | Acceleration Time | 0S |
| Pn016 | Time UPDN | 20 |
| Pn017 | Time BREAK | 10 |
| Pn018 | Cycles per loop | 3 |
| Pn019 | Loop WAIT | 1080S |
| Pn020 | Loop PER TEST | 20000 |

Table 2

| Pitch of screws shaft | Gear ratio | Magnetic poles | Actuator Type | Steps in each centimeter |
| --- | --- | --- | --- | --- |
| 5/2 | 40:1 | 3 | D, ES2, ER2 | 960 |
|  | 40:2 | 3 |  | 480 |
| 6/2 | 40:1 | 3 | ES3, ER3 | 400 |
|  | 40:2 | 3 |  | 200 |
| 7.5/3 | 40:1 | 3 | D | 640 |
|  | 40:2 | 3 |  | 320 |
| 16/4 | 40:1 | 3 |  | 300 |
|  | 40:2 | 3 |  | 150 |
| 10/4 | 44:1 | 2 | ET3 | 176 |
|  | 44:1 | 2 | ET2 | 352 |
|  | 39:2 | 2 | EF | 156 |
| 20/5 | 44:1 | 2 | ET2 | 176 |